Project name: GIGA Nuclear Magnetic Resonance – GIGA NMR

Number of the public contract: 302/7-O-CATS/26

Technical specifications for the “Purchase and delivery of the Equipment for the Center for Solid-State Analysis”

- SUBJECT OF PUBLIC TENDER

The subject of this public procurement is the purchase and supply of equipment for the Centre for Solid-State Analytics, comprising an integrated system for surface analysis, spatially resolved chemical characterization, and the determination of the chemical composition and chemical states of surfaces in demanding research and development environments. The equipment will enable correlative investigation of heterogeneous systems, thin films, functional surfaces, and trace species, and will constitute key infrastructure for advanced research and for supporting development in biotechnology, pharmacy, and advanced materials.

The public procurement is divided into five lots:

Lot 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS)
Lot  2: Near-atmospheric pressure X-ray photoelectron spectroscopy equipment (NAP-XPS)
Lot    3: Fully automated multi-analytical UHV system for the analysis of surfaces and interfaces
Lot     4: Equipment for three-dimensional surface characterization, optical imaging, and elemental

analysis of materials

Lot     5: System for nano- and micro infrared microscopy

These specifications relate to Lot 3: Fully automated multi-analytical UHV system for the analysis of surfaces and interfaces.

The subject of procurement for Lot 3 is the multi-analytical UHV system: supply, delivery, installation, commissioning, performance verification, site acceptance testing, and user training of a fully automated multi-analytical ultra-high vacuum surface and interface analysis system intended for comprehensive characterization of materials by complementary, integrated analytical techniques.

The system shall be designed to enable high-throughput and high-repeatability measurements under ultra-high vacuum conditions, with controlled and automated sample handling and transfer. It shall support advanced surface analysis workflows by combining multiple measurement modalities within a single, software-integrated platform, including (where included in the item schedule) techniques such as electron spectroscopies, ion scattering methods, electron energy-loss methods, optical spectroscopies (Raman under vacuum), and electron microscopy/microanalysis, together with depth profiling capabilities.

The procurement shall include a complete and operational system with all accessories, interfaces, and services listed in the item schedule, including integration of the analytical modules, necessary vacuum infrastructure, sample transfer solutions (including glovebox connectivity where required), control and data acquisition software, and all activities required to deliver a fully functional installation ready for routine multi-technique operation.

- TECHNICAL REQUIREMENTS

The minimum technical requirements for Lot 3 are set out below.

The tender must comply with or exceed all minimum technical requirements.

Delivery deadline: The tenderer (bidder)  must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 3.17 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 460 days from the entry into force of the contract.

Warranty period: For the delivered equipment, the bidder provides a 60 (sixty) months warranty for a flawless technical operation.

Lot 3: Fully automated multi-analytical UHV system for the analysis of surfaces and interfaces

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 3.1 | 1 | Surface analysis system base with integrated multi-energy monochromatic X-ray excitation that: • The bidder shall supply a fully automated surface analysis instrument base designed for high stability and high repeatability measurements, with preparedness for integration of the additional analytical modules listed in the item schedule. • The system shall include a fully enclosed main unit with a field-compensated analysis chamber and a fully motorized three-axis sample stage, including a bias connection enabling measurement of sample current. • The system shall support sample carriers of 50 mm × 20 mm or larger • The vacuum system shall be fully software controlled, include internal bakeout capability of the main analysis chamber and the introduction chamber, use oil-free rough vacuum pumping, and achieve a base pressure of 5·10–10 millibar or lower after bakeout. • The system shall include a preparation chamber enabling sample introduction without venting the main chamber and shall provide automated sample carrier transfer from the preparation chamber into the analysis chamber, with automated venting and pumping sequences and appropriate safety interlocks. The system shall include: - Turbomolecular pump on the analysis chamber with vibration isolation. - Titanium sublimation pump on the analysis chamber. - Turbomolecular pump(s) for differential pumping of the cluster ion source (where installed). - Turbomolecular pump on the UHV rear preparation chamber. - Turbomolecular pump on the introduction chamber. - High-capacity dry backing pump for the analysis chamber and first-stage pumping of the ion source (where installed). - Dry scroll-type backing pump for the UHV preparation chamber. - Dry scroll-type backing pump for the fast-entry/load-lock chamber. - All pumps shall be integrated into the system vacuum control architecture. The system shall include: - Separate mains input for system pumping and instrumentation. - Separate mains input dedicated to bakeout heating. - Independent electrical circuits for operational load and bakeout load. - Mains cables shall be supplied for both inputs. • The system shall include a sample introduction/transfer chamber incorporating at least three dedicated internal parking positions for sample holders The preparation chamber shall include: - Horizontal specimen probe for heating and/or cooling. - Compact linear shift mechanism (minimum 50 mm travel). - Compatibility with heating and cooling holders. • The system shall include a digital microscope (or equivalent imaging solution) enabling sample observation at the measurement position, navigation, and documentation. • The electron analyser shall be a hemispherical analyser with a wide-angle lens and a mean radius of 150 millimetres (or higher), equipped with a pulse-counting detector (or equivalent), with an energy resolution < 0.50 eV on Ag 3d5/2. • The system shall support angle-resolved XPS (ARXPS). • The analyser pass energy shall be continuously adjustable over a range that includes 1 to 400 eV (or higher than 400), under software control. • The analyser and detector system shall support photoelectron kinetic energy measurements compatible with X-ray excitation energies up to 5 keV (or higher), where such excitation is provided by the instrument. • The analyser lens/analyser settings and detector operation shall be fully motorized (where applicable) and software controlled, integrated into the main control environment for automated methods/recipes and traceable parameter logging. • The X-ray excitation shall be provided by a monochromator with a high-intensity, micro-focused source and a selectable X-ray spot size from 200 micrometres (or smaller) to 900 micrometres (or larger), with fully software-controlled operation. • The monochromated X-ray excitation shall provide, under software control, at least two selectable monochromated excitation lines including: - monochromated line for Al Kα at 1486 eV, and - monochromated line for Ag Lα at 2984 eV, - The monochromated source shall be fully motorized (where applicable) and integrated into the main control software for automated methods/recipes, including line selection, operating state, and beam/spot selection. • The system shall include automated anode switching between the available monochromated X-ray anodes. The supply shall include motorized control of anode position, motorized adjustment of anode-to-crystal angle, and motorized adjustment of crystal mirror settings. Full integration into the system control software for automated selection and operation is required. • The analysis area at the sample shall be selectable and shall cover a range from ≤ 200 µm up to ≥ 900 µm, measured at the sample plane (or equivalent definition), while maintaining stable monochromation and alignment. • The power of the X-rays at the largest spot size shall be: - ≥ 300 W for Al Kα - ≥ 300 W for Ag Lα • At site acceptance, the vendor shall demonstrate count rates. The system offered shall meet or exceed the following minimum benchmarks (or equivalent to better performance demonstrated under comparable acquisition conditions) on a pure silver sample: Maximum spot size at the sample surface (diameter); 900 µm; 300 µm; 150 µm; 900 µm. Reference FWHM target (eV) or lower; 1.0; 1.0; 1.0; 0.6. Minimum XPS count rate (cps) using Al Kα (1486 eV); 3.4106; 1.7106; 340103; 1.4106. • The system shall include dual-beam charge neutralization using low-energy electron and ion sources (or equivalent) suitable for analysis of insulating and charging samples. • The system shall include a sputter depth profiling capability using a monoatomic argon ion source, including a selectable scan area and an ion energy range of 500 eV (or lower) to 4 keV (or higher). • Sample heating: The system shall provide controlled sample heating to 1000 K (or higher) in vacuum with closed-loop control and temperature readback. • The system shall include a recirculating chiller compatible with the X-ray source and associated cooling circuits. The chiller shall operate on 200–230 V, 50/60 Hz, be fully compatible with the supplied system configuration, and be delivered as part of the complete installation. No external third-party cooling infrastructure shall be required beyond standard laboratory utilities. • The control and data system shall provide automated sample handling; sample visualization/management; integrated instrument control; method/recipe management with user-level permissions; quantitative evaluation and report generation; comprehensive logging for traceability and maintenance; and secure remote access for diagnostics and support. • The system shall include a dedicated rack area for system mains, programmable logic controller (PLC), and core electronics. |
| 3.2 | 1 | Ion scattering spectroscopy capability (ion scattering module) • The bidder shall supply an ion scattering spectroscopy capability fully compatible with the offered surface analysis system and integrated for routine operation within the same analysis chamber and control environment. • The capability shall include the required analyser operating mode for ion detection, including bipolar operation (or equivalent functionality) enabling ion scattering measurements in the required polarity modes. • The capability shall include an ion source enabling operation with helium as the working gas for ion scattering measurements. • The capability shall be fully integrated into the system software for method setup, execution, instrument control, and data acquisition, including automated coordination of analyser settings and ion source parameters. • The capability shall be delivered with all necessary hardware, interfaces, feedthroughs, vacuum-compatible components, and safety interlocks required for a complete, operational ion scattering measurement workflow. |
| 3.3 | 1 | Ultraviolet photoelectron spectroscopy capability (ultraviolet photoelectron module) • The bidder shall supply an ultraviolet photoelectron spectroscopy capability fully compatible with the offered surface analysis system and integrated for routine operation within the same analysis chamber and control environment. • The capability shall include a helium ultraviolet source providing at least helium I and helium II excitation lines, suitable for valence band and work-function measurements. • The ultraviolet source shall be fully compatible with the installed electron energy analyser and shall enable acquisition of ultraviolet photoelectron spectra using the same sample handling, positioning, and charge-management concepts provided with the system. • The capability shall be fully integrated into the system software for method setup, execution, instrument control, and data acquisition, including control of the ultraviolet source operating states and coordination with analyser settings. • The capability shall be delivered with all necessary hardware, interfaces, vacuum-compatible components, and safety interlocks required for a complete, operational valence photoelectron measurement workflow. |
| 3.4 | 1 | Cluster ion source for sputter depth profiling under ultra-high vacuum (cluster ion sputtering module) • The bidder shall supply a gas cluster ion source fully compatible with the offered surface analysis system for sputter depth profiling under ultra-high vacuum conditions. • The system shall be an argon cluster ion beam source with an acceleration energy up to 8 keV or higher. • The system shall provide selectable cluster sizes. • The supply shall include an ion source comprising a cluster generation unit and an ionisation chamber. • The system shall include an ion-optical column incorporating a Wien mass filter (or equivalent) for selection/control of the cluster ion species and beam composition. • The system shall include a differential pumping system suitable for maintaining ultra-high vacuum compatibility during operation. • The cluster ion source shall be fully integrated into the instrument control and data acquisition software, enabling automated sputter/depth profiling workflows, method control, and traceable logging of sputter parameters. |
| 3.5 | 1 | Scanning electron microscopy and scanning Auger microanalysis capability (electron microscopy and microanalysis module) • The bidder shall supply an electron microscopy and microanalysis capability fully compatible with the offered surface analysis system and integrated for routine operation within the same analysis chamber and control environment. • The module shall provide scanning electron microscopy of the sample surface for navigation, documentation, and selection of regions of interest for subsequent microanalysis. • The module shall provide scanning Auger microanalysis capability for spatially resolved elemental and chemical-state information, enabling acquisition of spectra and maps from user-defined regions of interest. • The system shall achieve a lateral (chemical/imaging) resolution of 0.1 µm or lower in SEM/SAM operating mode. • The module shall support correlation of image and chemical information for region-of-interest selection and pattern matching based on chemical contrast (or equivalent functionality enabling location of features using chemical information). • The primary electron energy shall be selectable over a range spanning from 1 keV (or lower) to 10 keV (or higher). • The electron beam current at the sample shall be up to at least 50 nA (or higher). • The module shall support positioning of analysis points/areas on the sample, coordinated with the system’s sample stage and imaging functions. • The module shall be fully software integrated for imaging, point analysis, line scans, and mapping, including storage of images. • The module shall include all necessary vacuum-compatible components, detectors, feedthroughs, electronics, and safety interlocks required for a complete and operational electron microscopy/microanalysis workflow. |
| 3.6 | 1 | Reflection electron energy loss spectroscopy (REELS) capability (reflection electron energy loss module) • The bidder shall supply a reflection electron energy loss spectroscopy capability fully compatible with the offered surface analysis system and integrated for routine operation within the same analysis chamber and control environment. • The capability shall enable acquisition of reflection electron energy loss spectra from solid surfaces under ultra-high vacuum, suitable for characterization of surface excitations and electronic structure features. • The REELS capability shall include an electron excitation source for bombardment of the sample surface with primary electrons with an upper energy limit of at least 1 keV, suitable for routine REELS operation. • The module shall be compatible with the installed electron energy analyser and shall provide the required electron excitation and detection geometry to perform reflection electron energy loss measurements on the sample positioned at the analysis location. • The capability shall be fully software integrated into the instrument control and data acquisition environment, enabling method setup, automated execution, and synchronized control of excitation/detection parameters and analyser settings. • The capability shall include all necessary vacuum-compatible hardware, power supplies, feedthroughs, electronics, and safety interlocks required for a complete and operational reflection electron energy loss measurement workflow. |
| 3.7 | 1 | Raman spectroscopy capability (Raman module) • The bidder shall supply a fully integrated Raman spectroscopy capability fully compatible with the offered surface analysis system and integrated for routine operation within the same platform. • The Raman module shall provide an integrated measurement position enabling Raman measurements on the sample at a defined and reproducible location coordinated with the system’s sample handling and positioning functions. • The Raman module shall be integrated in a manner that preserves the vacuum integrity and operational stability of the overall system during Raman measurements. • The Raman capability shall be fully integrated into the system software for method setup, execution, instrument control, and data acquisition, including storage of Raman datasets together with viewport filters and analysis chamber camera. • The supply shall include all necessary opto-mechanical interfaces, feedthroughs, vacuum-compatible components, alignment provisions, and safety interlocks required for a complete and operational Raman measurement workflow. • The Raman module shall include a laser source with a wavelength of 532 nm, suitable for Raman measurements, and integrated into the system for routine operation. • The Raman module shall be delivered with an integrated spectral reference library enabling automated spectral matching, identification, and reporting. The library shall be fully accessible from the Raman control and data analysis software and shall support import of user spectra, local storage, and version control or equivalent management functions. - The proposed solution shall include integrated Raman spectral libraries with a total content of at least 26,000 spectra. The libraries shall provide broad coverage of compounds relevant to chemical, pharmaceutical, polymer, environmental, laboratory, and security applications, including pharmaceuticals, narcotics, excipients, polymers, household and industrial chemicals, pesticides, and laboratory reagents. |
| 3.8 | 1 | Glovebox system A glovebox system shall be provided for controlled sample handling and transfer to the main analysis chamber system for surface analysis. The proposed solution shall be fully compatible with the offered instrument configuration and shall enable safe, efficient, and contamination-controlled transfer of samples between the glovebox and the analytical chamber. Minimum required specifications The glovebox system shall include at least the following: • Metal-based enclosure with a large front viewing window and a suitable support stand. • An integrated mini antechamber for loading samples into the glovebox without exposing the internal glovebox atmosphere to ambient air or external gases. • The mini antechamber shall be sufficiently sized to accommodate loading of sealed sample transfer accessories of appropriate dimensions, including a capsule-type transfer container and a short transfer rod or equivalent. • A rear-wall vacuum-compatible flange connection, for direct connection to the extended entry/transfer lock of the analytical system. • An extended secondary transfer chamber (pre-lock or equivalent) with two loading ports. One door shall provide direct access to the chamber, while the second door shall provide access through the glovebox. • An extended transfer arm or equivalent manipulation system enabling sample loading and handling from both loading positions to the analytical chamber. • A single-column recirculating gas purification system, suitable for operation on standard mains supply, vacuum pump, manual solvent removal provision, and automatic purge valve. • Integrated oxygen and moisture monitoring sensors with a measurement range covering at least 0 to 1000 ppm. • Provision for variable feedthrough configurations, to be defined during the final technical design and integration stage. |
| 3.9 | 1 | Sample holders (sample holder and mounting set) • The bidder shall supply a set of sample holders and mounting fixtures fully compatible with the offered system’s sample stage, introduction chamber transfer concept, and all specified analytical techniques. • All holders shall be suitable for operation under ultra-high vacuum and designed to withstand the system bakeout conditions without loss of functionality. • The package shall include five sample holders, suitable for measurements and enabling secure, repeatable mounting of the system’s standard sample carrier format (50 mm × 20 mm or larger), including reliable electrical contact where required (e.g., biasing and sample current measurement). • The package shall include one sample holder for heating at least up to 1000 K or higher, supplied together with a dedicated power supply and the necessary feedthroughs/cabling for operation in vacuum. The heating holder shall support controlled heating for in-situ experiments and shall include a temperature sensing solution (e.g., thermocouple or equivalent) suitable for closed-loop control. • The package shall include one additional sample holder for heating from -150 °C (or lower) up to 300 °C (or higher) • The system shall be supplied with at least one separate sample holder suitable for powder samples. • An additional sample holder for thicker specimens shall be included, enabling analysis of samples with heights of at least 10 mm or greater. • The package shall include one sample holder for continuous sample rotation, enabling controlled azimuthal rotation during measurements (motorized or equivalent), with full compatibility with the system’s positioning and analysis geometry. • The supply shall include all necessary adapters, clamps, fasteners, and alignment aids required for routine sample exchange, safe transfer through the introduction chamber, and reproducible repositioning during multi-technique workflows. |
| 3.10 | 1 | Inverse photoelectron spectroscopy capability (inverse photoelectron module) • The bidder shall supply an inverse photoelectron spectroscopy capability fully compatible with the offered surface analysis system and integrated for routine operation within the same analysis chamber and control environment. • The capability shall enable measurement of unoccupied electronic states and related parameters (e.g., conduction-band features and work-function–related information) on solid surfaces under vacuum conditions. • The module shall provide a complete excitation and detection concept suitable for inverse photoelectron spectroscopy, including an electron beam source (or equivalent excitation) and a photon detection unit (or equivalent) optimized for inverse photoelectron measurements. • The module shall support stable and reproducible alignment of the excitation/detection geometry at the analysis position and shall be compatible with the system’s sample positioning functions, including repeatable return to predefined measurement points. • The capability shall be fully integrated into the system software for method setup, execution, instrument control, and synchronized data acquisition, including storage of inverse photoelectron datasets. • The supply shall include all required vacuum-compatible hardware, power supplies, feedthroughs, electronics, and safety interlocks necessary for a complete and operational inverse photoelectron spectroscopy workflow. |
| 3.11 | 1 | Licensed software for measuring and processing data for at least 10 users. The control and analysis software shall provide full instrument control and data processing, including background subtraction, peak fitting, quantification, and automated reporting, for UHV operation. |
| 3.12 | 1 | Dedicated process computer with the following specifications:* • MS Windows 11 Professional 64-bit operating system, • desktop/workstation computer with a processor with at least 6 physical cores and a maximum boost frequency of at least 5.7 GHz, or better, • 32 GB RAM or more, 1 TB SSD or more, • 1 × 27" LED display (27" or larger) with Full HD resolution or higher than Full HD resolution. * Special requirements for item 3.12 in accordance with the Decree on Green Public Procurement are set out at the end of this document. |
| 3.13 | 1 | Warranty (minimum 60 months coverage). The bidder shall provide a minimum total coverage period of 60 months from successful installation and site acceptance test. • It shall cover replacement of defective parts and all corresponding labour costs. • It shall cover travel and living expenses of the bidder’s personnel for any required on-site intervention. • It shall cover all spare parts required to maintain the instrument in optimum condition and high availability, at no additional cost during the covered period. • The package shall cover all consumables required according to lifetime/condition, including on-site replacement/exchange performed by the bidder’s expert personnel. • It shall include at least one preventive maintenance visit per year with a minimum on-site duration equivalent to two working days (8 hours per working day), including all travel and accommodation costs. • It shall include unlimited corrective visits as required, including emergency visits, with labour, travel, accommodation, and necessary parts included. • It shall include calibration, testing, and performance verification activities required to keep the instrument within specification. • It shall include unlimited remote diagnostics and support (e-mail and secure remote access/remote desktop). • The package shall include free software updates for the installed control and data analysis software. • The package shall include extended customer training consistent with the service offered level. • The bidder shall provide on-site service within one week and respond to email notifications within 48 hours. |
| 3.14 | 1 | Basic operation training • Acquaintance with the equipment and basic training, performed by the engineer of the bidder in Slovene or English, for up to four users lasting at least 5 working days; 8 hours per working day (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the bidder’s engineer; training is carried out at the location of FKKT UM, determined by the client). The item is completed after the delivery, installation, and startup of the equipment. |
| 3.15 | 1 | The advanced user training package shall include two separate advanced on-site training sessions after acceptance. The training shall be conducted at the FKKT UM location designated by the contracting authority, in Slovene or English. Each advanced training session shall last 4 working days, with 8 hours of training per working day, and shall be no later than 31 days after completion of the basic training referred to in item 3.14. Advanced training shall be provided for a maximum of four persons. The contractor shall bear all travel, subsistence, and accommodation costs of the personnel delivering the training. |
| 3.16 | 1 | Provided remote access to the device for the needs of calibration, repair, and control of the device, at least for the duration of the warranty period (60 months). The instrument design must facilitate remote operation and fault diagnostics. |
| 3.17 | 1 | Delivery and installation on the client’s premises; 3rd floor at FKKT UM. Conditions for successful completion of the item and signing of the acceptance report: - delivery, installation, and commissioning of the equipment at the contracting authority’s site, to be carried out by the bidder, - familiarization with the equipment and basic training under item 3.14, - handover of the operating instructions and other technical documentation in digital form and in the English language. The performance specifications (items 3.1 to 3.14) must, prior to the signing of the acceptance documents, be demonstrably fulfilled on-site (FKKT UM). |

SPECIAL REQUIREMENTS FOR ITEM 3.12

The equipment under item 3.12, Dedicated Process Computer, must fully comply with the technical specifications, ensuring the fulfillment of the objectives in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 121/21, 132/23, and 43/25), namely:

DEDICATED PROCESS COMPUTER

Power Consumption MANDATORY: The personal computer must belong to the highest energy efficiency class available on the market. Method of proof: The bidder must attach the following to the tender: - a certificate showing that the product has a Type I environmental label or a label from another certification scheme (e.g. Energy Star V9.0, EPEAT, TCO) that meets the specified requirements. Appropriate supporting documents may also include alternative test results obtained by testing bodies accredited in accordance with ISO 17025, in compliance with IEC 62623.

ELECTRONIC DISPLAYS

Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market, an energy class of at least B,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL).

The operation is co-financed by the European Union from the European Regional Development Fund under the European Cohesion Policy Programme 2021–2027, Policy Objective 1: A more competitive and smarter Europe by promoting innovative and smart economic transformation and regional ICT connectivity, under Priority 11: Development of Strategic Technologies for Europe – STEP, and under Specific Objective RSO1.6: Support for investments contributing to the objectives of the Strategic Technologies for Europe Platform (STEP).