Project name: GIGA Nuclear Magnetic Resonance – GIGA NMR

Number of the public contract: 302/7-O-CATS/26

Technical specifications for the “Purchase and delivery of the Equipment for the Center for Solid-State Analysis”

- SUBJECT OF PUBLIC TENDER

The subject of this public procurement is the purchase and supply of equipment for the Center for Solid-State Analytics, comprising an integrated system for surface analysis, spatially resolved chemical characterization, and the determination of the chemical composition and chemical states of surfaces in demanding research and development environments. The equipment will enable correlative investigation of heterogeneous systems, thin films, functional surfaces, and trace species, and will constitute key infrastructure for advanced research and for supporting development in biotechnology, pharmacy, and advanced materials.

The public procurement is divided into five lots:

Lot 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS)
Lot 2: Near-atmospheric pressure X-ray photoelectron spectroscopy equipment (NAP-XPS)
Lot 3: Fully automated multi-analytical UHV system for the analysis of surfaces and interfaces
Lot 4: Equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials

Lot 5: System for nano- and micro infrared microscopy

These specifications relate to Lot 4: Equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials

The subject of procurement for Lot 4 is equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials: The system shall enable non-contact three-dimensional surface characterization, high-resolution optical imaging, and rapid elemental identification of solid materials, supporting surface topography measurements, microstructural inspection, and localized elemental analysis of materials such as metals, alloys, coatings, semiconductors, ceramics, polymers, and composites. The procurement shall include a 3D laser scanning confocal microscope with white-light interferometry, a digital optical microscope with motorized positioning, and a laser-based elemental analyzer based on laser-induced breakdown spectroscopy. The delivered system shall be fully operational, including all required components, software, installation, commissioning, performance verification, acceptance, and user training.

- TECHNICAL REQUIREMENTS

Delivery deadline: The tenderer (bidder) must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 4.4 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 30 days from the entry into force of the contract.

Warranty period: For the delivered equipment, the bidder provides a 24 months warranty for a flawless technical operation.

Lot 4: Equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 4.1 MAIN EQUIPMENT | 1 | 3D Laser Scanning Confocal Microscope System: Measurement system The system shall be based on laser scanning confocal microscopy combined with white-light interferometry and focus variation in one system for 3D surface characterization. Minimum requirements: - Confocal laser scanning measurement capability - White-light interferometry measurement capability - Automatic switching between measurement modes - Non-contact 3D surface measurement - Capability for large-area surface stitching - Capability for surface roughness and step-height measurement - Capability for measurement of reflective and non-reflective surfaces Measurement head The system shall include a confocal laser measurement head. Minimum requirements: - Laser wavelength in the red range (661 nm) - High-precision vertical scanning system - Integrated optical system for confocal measurement - Vertical measurement display resolution ≤ 1 nm - Compatible with the interferometry module System controller The system shall include a dedicated microscope controller unit for instrument operation and data acquisition. Minimum requirements: - Control of the laser scanning system - Control of stage movement and focusing - Data acquisition and signal processing - Interface for external control computer - High-speed data communication with control PC Motorized stage system The microscope shall include a motorized stage for automated sample positioning. Minimum requirements: - Motorized X-Y stage for automated positioning - Motorized Z-axis focusing system - Support for sample stitching measurements - Motorized X–Y stage travel ≥ 100 × 100 mm Tilting stage The system shall include a tilting sample stage. Minimum requirements: - Adjustable sample tilt - Stable positioning during measurement - Compatibility with the motorized stage system Interferometry module The system shall include a white-light interferometry module for ultra-high-resolution height measurements. Minimum requirements: - Spectral interferometry measurement principle - Interferometric height display resolution ≤ 0.1 nm - Automatic integration with the confocal measurement system Stitching module The system shall include a software or hardware module enabling automatic stitching of multiple measurement areas. Minimum requirements: - Automated acquisition of tiled measurements - Automatic reconstruction of large measurement areas - Stitching of multiple fields of view into a single 3D dataset Optical components The system shall include at least the following optical components: - High-magnification objective lens - Plan apochromatic objective - Magnification 150×, 50×, 20×, 10×, 5× - Interference objective lens - Magnification 10× - Optimized for interferometric measurements |
| 4.2 MAIN EQUIPMENT | 1 | Digital Optical Microscope System The bidder shall provide a digital optical microscope system for high-resolution imaging, documentation, and automated inspection of microstructures, material surfaces, and material analysis. The system shall include all hardware, optics, control electronics, software, and accessories required for full operation. Digital microscope controller The system shall include a digital microscope controller unit. Minimum requirements: - Compatible with high-resolution digital microscope cameras - Progressive scanning imaging system - Image sensor resolution at least 2064 × 1536 pixels - Image processing with 16-bit dynamic range - Adjustable electronic shutter with ≤ 1/19,000 s - Automatic and manual gain control Digital microscope camera The system shall include a high-performance digital microscope camera. Minimum requirements: - Total pixel count 3.15 megapixels or higher - Resolution modes at least 2064 × 1536 pixels - Progressive image acquisition - High dynamic range imaging - Automatic and manual exposure control - Automatic and manual white balance Motorized microscope stand The system shall include a motorized free-angle microscope stand with automated positioning and recognition of tilt angle. Minimum requirements: - Motorized Z-axis focusing - Motorized X-Y stage movement - X-Y stage travel at least 40 mm - Z-axis measurement range at least 45 mm - Stage rotation capability up to ±90° Zoom lens The system shall include a high-performance zoom lens. Minimum requirements: - Magnification ranges from 20× to 200× - Optical zoom with adjustable field of view - Working distance 25 mm or higher - High depth of field at least 33 mm Lens mounting system The system shall include: - Lens joint compatible with the zoom lens - Lens holder for stable optical alignment Display and control console The system shall include a dedicated console for handling the X-, Y-, and Z-axis Minimum requirements: - Adjustable viewing angle - Integrated system control interface |
| 4.3 SUPPORTING EQUIPMENT | 1 | Laser-Based Elemental Analyzer The system shall be based on laser-induced breakdown spectroscopy (LIBS). Minimum requirements: - Laser ablation with spectroscopic plasma analysis - Broadband optical emission spectroscopy detection - Capability for rapid multi-element detection Laser system The system shall include an integrated analytical laser source. Minimum requirements: - Nd:YAG laser - Laser wavelength 355 nm - Pulse energy sufficient for plasma generation - Focused laser spot size 30 µm or lower Spectroscopic detection system The system shall include a spectrometer for elemental emission analysis. Minimum requirements: - Simultaneous multi-element detection capability - Optical emission spectrum acquisition from plasma emission Observation system The analyzer shall include an integrated optical observation system. Minimum requirements: - Observation magnification range 300x–1000x or wider - Progressive image acquisition - Real-time observation of analysis location Imaging sensor Minimum requirements: - Progressive image acquisition - Frame rate 50 fps or higher Optical system Minimum requirements: - Objective lens working distance 25 mm or higher - Optical configuration optimized for laser focusing and plasma observation Image acquisition and illumination Minimum requirements: - Adjustable illumination intensity |
| 4.4 SUPPORTING EQUIPMENT | 1 | Control software, computing hardware, and training The bidder shall provide software, computing hardware, and user training required for operation, control, data acquisition, and analysis of the microscopy and analytical systems without any additional costs. Control and analysis software The system shall include integrated software for instrument control and data analysis. Minimum requirements: - Control of microscope systems and analytical modules - Image acquisition and data recording - 3D surface reconstruction and visualization - Surface roughness and profile analysis - Measurement stitching and large-area reconstruction - Image measurement and annotation tools - Data export in standard formats in TIFF, JPEG, and CSV Control computer and monitor* The system shall include a dedicated desktop computer for system operation. Minimum requirements: - Preconfigured for instrument control - Compatible with instrument software - MS Windows 11 Professional 64-bit operating system, - Desktop/workstation with i5-11600 processor or better, 32 GB RAM or higher, 1 TB SSD hard disk or higher, 27-inch (or larger) LED monitor with Full HD resolution or with higher resolution than Full HD. Control computer and monitor The system shall include a dedicated desktop computer for system operation. Minimum requirements: - Preconfigured for microscope control - Compatible with instrument software - MS Windows 11 Professional 64-bit operating system, - Desktop/workstation with i5-11600 processor or better, 16 GB RAM or higher, 1 TB SSD hard disk or higher, 23-inch (or larger) LED monitor with Full HD resolution or with higher resolution than Full HD. Measurement and analysis software The system shall include software for acquisition, visualization, and analysis of 3D surface data. Minimum capabilities: - Instrument control and measurement automation - 3D surface reconstruction - Surface roughness analysis - Step-height measurement - Surface profile extraction - Measurement stitching - Surface recommendation function depending on the surface roughness of different samples Data export and reporting Software The system shall include image acquisition and analysis software. Minimum capabilities: - Real-time image acquisition - Image measurement tools - Depth-of-field synthesis - Image stitching - Data export and reporting Installation and commissioning The supplier shall provide installation and system commissioning, including: - Hardware installation - Software configuration - Functional verification of all system components User training The supplier shall provide operator training for system users. Minimum requirements: - Training in instrument operation in Slovene or English, measurement procedures, and data analysis and reporting for at least 4 persons on-site after installation, which must include 2 sessions of 4 hours each on 2 consecutive days. Technical support and software updates The supplier shall provide for at least 3 years: - Technical support for system operation - Software updates for the control and analysis software - Maintenance support. *Special requirements for the process computer and display under item 4.4 in accordance with the Decree on Green Public Procurement are set out at the end of this document. |
| 4.5 | 1 | Warranty (minimum 24 months total coverage). The bidder shall provide a minimum total coverage period of 24 months (2 years) from successful installation and site acceptance test. - It shall cover replacement of defective parts and all corresponding labour costs. - It shall cover travel and accommodation expenses of the supplier’s personnel for any required on-site intervention. - It shall cover all spare parts required to maintain the instrument in optimum condition and high availability, at no additional cost during the covered period. - The service package shall cover all consumables required according to lifetime/condition, including on-site replacement/exchange performed by the supplier’s expert personnel. - It shall include unlimited corrective service visits as required, including emergency visits, with labour, travel, accommodation, and necessary parts included. - It shall include calibration, testing, and performance verification activities required to keep the instrument within specification. - It shall include unlimited remote diagnostics and support (e-mail and secure remote access/remote desktop). - The service package shall include free software updates for the installed control and data analysis software. The service package shall include extended customer training consistent with the offered service level. |
| 4.6 | 1 | Delivery and installation on the client’s premises (DDP FKKT UM, Maribor); 3rd floor at FKKT UM. Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation, and commissioning of the equipment on the client’s premises by the supplier, - familiarization with the equipment and basic training under item 4.4, - submission of the operating instructions in digital form in Slovene or English, and of other technical documentation, - performance specifications (items from 4.1 to 4.4) must be demonstrated on-site (FKKT UM) before signing acceptance documents. |

SPECIAL REQUIREMENTS FOR ITEM 4.4

The equipment under item 4.4, Dedicated Process Computer, must fully comply with the technical specifications, ensuring the fulfillment of the objectives in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 121/21, 132/23, and 43/25), namely:

DEDICATED PROCESS COMPUTER

Power Consumption MANDATORY: The personal computer must belong to the highest energy efficiency class available on the market. Method of proof: The bidder must attach the following to the tender: - a certificate showing that the product has a Type I environmental label or a label from another certification scheme (e.g. Energy Star V9.0, EPEAT, TCO) that meets the specified requirements. Appropriate supporting documents may also include alternative test results obtained by testing bodies accredited in accordance with ISO 17025, in compliance with IEC 62623.

ELECTRONIC DISPLAY (27-inch (or larger) LED monitor with Full HD resolution or with higher resolution than Full HD.

Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market of at least B,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL).

ELECTRONIC DISPLAY (23-inch (or larger) LED monitor with Full HD resolution or with higher resolution than Full HD.

Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market A,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL).

The operation is co-financed by the European Union from the European Regional Development Fund under the European Cohesion Policy Programme 2021–2027, Policy Objective 1: A more competitive and smarter Europe by promoting innovative and smart economic transformation and regional ICT connectivity, under Priority 11: Development of Strategic Technologies for Europe – STEP, and under Specific Objective RSO1.6: Support for investments contributing to the objectives of the Strategic Technologies for Europe Platform (STEP).