Project name: GIGA Nuclear Magnetic Resonance – GIGA NMR

Number of the public contract: 302/7-O-CATS/26

Technical specifications for the “Purchase and delivery of the Equipment for the Center for Solid-State Analysis”

- SUBJECT OF PUBLIC TENDER

The subject of the public procurement is the purchase and delivery of equipment for the Center for Solid-State Analytics, constituting an integrated system for surface analysis, spatially resolved chemical characterization, and the determination of the chemical composition and chemical states of surfaces in demanding research and development environments. The equipment will enable the correlative investigation of heterogeneous systems, thin films, functional surfaces, and trace species, and will provide key infrastructure for advanced research and support for development in biotechnology, pharmaceuticals, and advanced materials.

The public procurement is divided into five lots:

Lot 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS)

Lot 2: X-ray photoelectron spectroscopy equipment for operation at near-ambient pressure (NAP-XPS)

Lot 3: Fully automated multi-analysis UHV system for the analysis of surfaces and interfacial layers

Lot 4: Equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials

Lot 5: System for nano- and micro-infrared microscopy

These specifications relate to Lot 5: System for nano- and micro-infrared microscopy

The subject of procurement under Lot 5 is a nano- and micro-infrared microscopy system, including supply, delivery, installation, commissioning, performance verification, acceptance, and user training for an advanced system intended for spatially resolved chemical and structural characterization of surfaces and materials at the micro- and nanometer scale. The system must enable nano-infrared imaging and spectroscopy with very high spatial resolution below the diffraction limit of conventional infrared microscopy and must be suitable for demanding sample types, including thin films, multiphase materials, functional surfaces, coatings, polymeric and inorganic systems, corrosion products, and other heterogeneous samples. The subject of procurement under this lot is the supply of a complete, fully operational nano-infrared microscopy system, together with all additional equipment and services specified in the item table below, and compliance with the detailed technical requirements for each listed item.

- TECHNICAL REQUIREMENTS

The minimum technical requirements for Lot 5 are set out below.
The offer must meet or exceed all minimum technical requirements.

Delivery deadline: The tenderer (bidder) must supply and deliver the equipment that is the subject of this public procurement/lot to the University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, and must carry out there all activities required for the successful completion of item 5.22 of the technical specifications no later than 270 days from the effective date of the contract.

Warranty period: For the delivered equipment, the bidder must provide a 24 (twenty-four) month warranty for faultless technical operation.

Lot 5: System for nano- and micro-infrared microscopy

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 5.1 MAIN EQUIPMENT | 1 | Integrated platform for near-field optical microscopy with scanning probe microscopy for nano-infrared imaging and spectroscopy The bidder shall offer one integrated research platform for near-field optical microscopy, coupled with a scanning probe microscopy sensor (SPM/AFM) or an equivalent high-resolution method, suitable for nano-infrared mapping, point spectroscopy, and correlative topographic characterization of samples at the micro- and nanometer scale. The offered system must meet at least the following requirements: - The system must be intended for nano-optical measurements and must enable correlation of topographic data with the local optical or infrared responses of the sample. - The system must enable at least: - measurement of surface topography, - selection and precise positioning of the region of interest, - spatially resolved nano-optical or nano-IR mapping, - point measurements or spectroscopic measurements at the selected location on the sample. - The system must be compatible with optical excitation in the visible and infrared spectral ranges. Broader spectral capabilities are also acceptable if the supplier demonstrates equivalent or better research applicability. - An AFM for nanoscale characterization suitable for optical/nano-IR mapping and spectroscopy must be included. - The system must include motorized sample movement in the X, Y, and Z axes for sample loading and selection of the measurement location. The minimum travel ranges shall be: - at least 55 mm in the X-direction, - at least 15 mm in the Y-direction, - at least 5 mm in the Z-direction. - The step size in all three axes must be less than 200 nm. - The system must include an XY scanner for fine sample scanning with a closed-loop range of at least 100 µm × 100 µm. - The resolution of fine XY scanning must be in the sub-nanometer range. - The system must include a Z scanner for fine height control with effective vertical resolution in the sub-nanometer range. The minimum required measurement range in the Z-axis shall be at least 2.5 µm. - The system must allow operation with samples of at least 30 mm × 40 mm × 15 mm, or with an equivalent configuration providing comparable or better research applicability. - The measuring head must be compatible with standard AFM cantilevers or equivalent probes for nanoscale characterization. Operation with probes having a resonance frequency of at least up to 500 kHz, or equivalent, must be supported. - The measuring head must enable motorized alignment in the X, Y, and Z axes for adjustment of the probe position relative to the optical beam or measurement geometry. - The system must include an optical microscope for sample inspection and locating the measurement position, with a resolution better than 1 µm, as well as a digital camera for real-time sample visualization. - The system must support at least contact-mode and intermittent contact-mode AFM measurements. - The system must include an optical assembly for focusing and signal collection in the region of the measuring probe for nano-optical and nano-IR measurements. - The system must enable motorized and repeatable adjustment of the optical components for focusing on the measurement region. - The control electronics must enable synchronization of mechanical and optical signals and connection of external signals. The system must include at least: - two analog input channels or equivalent, - two analog output channels or equivalent, - two digital I/O channels or equivalent. - The system must include control, acquisition, and analysis software for AFM and optical measurements. The software must enable: - instrument control, - 1D, 2D, and 3D data acquisition, - real-time signal visualization. - The scanning probe head of the system must accept AFM cantilevers with a resonance frequency up to 500 kHz or higher. - The absolute position of the AFM probe and AFM head must remain unchanged during sample exchange. |
| 5.2 MAIN EQUIPMENT | 1 | High numerical aperture optical objective for near-field measurements, nano-IR mapping and spectroscopy The bidder shall offer an optical objective or an equivalent high numerical aperture optical assembly intended for focusing the excitation radiation onto the region of the measurement tip and for efficient collection of the optical signal in near-field measurements, nano-IR mapping, and spectroscopy. The offered solution must meet at least the following requirements: - the optical assembly must be suitable for near-field measurements and for nano-IR mapping and spectroscopy; - the optical assembly must enable the use of two simultaneous light beams in the region of the measurement tip or equivalent functionality; - the numerical aperture of the optical assembly must be at least NA 0.45; - the optical assembly must be compatible at least with illumination in the visible and infrared spectral ranges and must also be compatible with the THz range; - the optical assembly must ensure an improved signal-to-noise ratio in near-field measurements; - the optical assembly must be compatible with the offered nano-IR microscopy and spectroscopy system; - in order to ensure efficient focusing of light onto the AFM tip, the parabolic mirror objective must allow translational movement relative to the optical axis of the incident illumination sources; - it must be possible to store and recall predefined objective positions in 3D space, thereby ensuring a default instrument alignment state; - the system must support two independent optical beam paths: for illumination of the AFM tip and for detection of scattered light, to support near-field imaging and spectroscopy, as well as THz TDS spectroscopy; - the integrated parabolic mirror must include a motorized positioning assembly for precise focusing of an external light source onto the AFM tip. This positioning assembly must have a travel range of 4 mm in X, Y, and Z and must be fully controlled via the user software in order to enable focusing of light sources in different spectral ranges (visible to IR and THz) and under different collimation conditions; - the system must allow an upgrade with a THz-TDS light source to enable future nano-THz measurements. |
| 5.3 MAIN EQUIPMENT | 1 | High z-Range Sample Scanner The bidder shall offer a sample scanner with extended travel in the Z-direction, compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable extension of the sample scanning range in the Z-direction to at least 7 µm; - it must be suitable for measurements of samples with an inclined surface or with larger height differences; - it must be compatible with the offered system. |
| 5.4 MAIN EQUIPMENT | 1 | Active piezo vibration isolation workstation The bidder shall offer an active vibration isolation workstation compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must be based on active piezoelectric vibration isolation; - it must provide vibration isolation of at least 25 dB at 5 Hz; - it must provide vibration isolation of at least 35 dB at frequencies above 10 Hz; - the load capacity must be at least 150 kg; - it must include automatic load leveling; - it must include an optical breadboard with dimensions of at least 150 × 100 cm; - it must include a supporting metal frame; - the external dimensions of the system must not exceed 180 × 130 × 80 cm (W × D × H); - it must be compatible with the offered system. |
| 5.5 MAIN EQUIPMENT | 1 | High-Precision Synchronization I/O Board The bidder shall offer a high-precision synchronization input/output board compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must include at least 2 high-precision synchronized output channels in a frequency range of at least 1 kHz to 1 MHz; - it must include at least 2 high-precision synchronized input channels; - it must enable simultaneous near-field correlation photocurrent mapping or equivalent functionality; - it must be compatible with the offered system; - the system must support simultaneous demodulation and acquisition at each pixel from the 1st to the 5th harmonic of the tip frequency, for at least 4 independent signals (mechanical/AFM, s-SNOM signals, each with amplitude and phase, and 2 additional channels). |
| 5.6 | 1 | Software for amplitude- and phase-resolved near-field mapping The bidder shall offer software compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable amplitude- and phase-resolved near-field mapping; - it must enable high-performance near-field mapping; - it must provide spectral resolution of up to 1 cm⁻¹ or better than 1 cm⁻¹; - it must be compatible with the appropriate detection module; - it must be compatible with the offered system. |
| 5.7 MAIN EQUIPEMENT | 1 | Point Spectroscopy Upgrade The bidder shall offer a point spectroscopy upgrade compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable the acquisition of amplitude- and phase-resolved near-field spectra; - it must provide spectral resolution of up to 1 cm⁻¹ or better than 1 cm⁻¹; - it must be compatible with the mid-IR illumination unit; - it must be compatible with the detection module in the white-light position; - it must be compatible with the near-field imaging unit; - it must be compatible with the offered system; - the system must provide near-field spectroscopy in the mid-IR range with continuous sweeping of the IR laser illumination. This spectroscopic modality must be based on a complete background-free near-field optical detection technology using interferometric detection and must result in amplitude- and phase-resolved spectra; - the system must enable spectroscopic acquisition of both s-SNOM amplitude and phase using wavelength-tunable sources, without the need for additional reference measurements. |
| 5.8 | 1 | Spectroscopy in tapping AFM-IR mode and software for absorption nano-IR mapping The bidder shall offer software compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable absorption nano-IR mapping and spectroscopy; - it must enable operation in tapping AFM-IR mode; - it must be compatible with the appropriate illumination unit; - it must be compatible with the offered system. |
| 5.9 MAIN EQUIPMENT | 1 | White-Light Reflection-Mode Detection Module for VIS, NIR and MIR, motorized version The bidder shall offer a white-light reflection-mode detection module, in motorized version, compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable detection for high-performance near-field mapping; - it must enable simultaneous detection of the amplitude and phase of the optical near-field signal; - it must enable near-field mapping and spectroscopy in the white-light position; - it must enable point spectroscopy by laser sweeping or equivalent functionality; - it must be suitable for detection in the VIS, NIR and MIR ranges; - it must be fully motorized and software-controlled; - it must enable an AFM scan speed of at least up to 20 µm/s at the highest spatial resolution; - it must enable reflective AFM-tip illumination; - it must be compatible with the offered system; - the optical detection module must support interferometric detection of the near-field signal for amplitude- and phase-resolved measurements, with a minimum integration time of 4 ms/pixel; - the module must enable detection of the near-field signal with complete background suppression, including multiplicative background. The nanoscale near-field confinement of the optical interaction between the tip and the sample at higher harmonics of the optical signal, as determined by vertical scans (approach curves), must be less than 20 nm; - the detection module must operate in reflection mode (side illumination of the tip for opaque samples) and must be suitable for wavelengths in the visible, near-infrared and mid-infrared ranges; - the module must support an ultra-compact optical beam path design such that the interferometric arm length is < 10 cm, which is required for stable near-field measurements; - the module must include an exchangeable beam-splitter holder mount for replacing beam splitters for measurements in the visible and IR ranges; the mounting system must include a ball-and-V-groove design for high-precision positioning and easy switching of the illumination unit. |
| 5.10 MAIN EQUIPMENT | 1 | Near-Field Imaging Unit The bidder shall offer a near-field imaging unit compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must be optimized for high-performance near-field imaging; - it must enable coupling to the microscope system with motorized beampath control; - it must include at least 3 separate optical inputs for VIS, NIR, and MIR; - it must include a powermeter for laser power calibration; - it must include at least 1 optical output for an external detector; - it must enable software-controlled attenuation, polarization, and detector alignment; - it must be compatible with the offered system. |
| 5.11 MAIN EQUIPMENT | 1 | Near-Field Imaging Unit with MIR Detection The bidder shall offer an upgrade of the near-field imaging unit with MIR detection, compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must enable MIR detection of tip-scattered light; - it must enable automated detector selection with compatible MIR illumination units; - it must enable motorized power control of the MIR beam; - it must enable software-controlled alignment of the MIR detector; - the detector must be optimized for measurements at least in the range from 5 µm to 15 µm; - the detector sensitivity must be at least from 3 µm to 16 µm; - it must be compatible with the offered system. |
| 5.12 MAIN EQUIPEMENT | 1 | Widely tunable OPO laser source The bidder shall offer a widely tunable OPO laser source compatible with the offered nano-IR microscopy and spectroscopy system. The offered solution must meet at least the following requirements: - it must be suitable for near-field mapping and for mapping and spectroscopy; - the tuning range must cover at least 1.4 µm to 16 µm; - in the mid-IR range, it must cover at least 650 cm⁻¹ to 2000 cm⁻¹; - in the idler range, it must cover at least 2400 cm⁻¹ to 4400 cm⁻¹; - in the signal range, it must cover at least 5000 cm⁻¹ to 6800 cm⁻¹; - the output power must be at least 150 mW at 3800 nm; - the output power must be at least 2 mW at 12000 nm; - the pulse duration must be approximately 10 ps or less; - the spectral linewidth must be less than 4 cm⁻¹; - the repetition rate must be at least 50 MHz; - it must enable motorized tuning across the entire spectral range; - it must be compatible with the offered system; - the OPO laser system must have a repetition rate in the range of 80–90 MHz, a pulse duration of < 10 ps, and a spectral linewidth of < 5 cm⁻¹ at an average output power of > 8 mW at a wavelength of 12 µm; the OPO emission must be dynamically controllable in pulsed mode down to the kHz range, with automatic control via software; - the OPO illumination unit must be compatible with AFM-IR and with s-SNOM imaging and spectroscopy. |
| 5.13 MAIN EQUIPMENT | 1 | Additional AFM modes The bidder shall offer additional AFM modes compatible with the offered system for nanoscale characterization of surfaces and local material properties. The offered solution must meet at least the following requirements: - it must include at least the following AFM modes: KPFM, Piezo Force Microscopy, Tip Force Microscopy, and conductive AFM, or equivalent modes; - it must enable measurement of local surface potential, piezoelectric response, tip forces, and local electrical conductivity; - it must be software-compatible and functionally compatible with the offered AFM system; - it must allow use within the same measurement system without the need to replace the basic platform; - it must include all necessary software and hardware support for operation of the stated AFM modes; - Kelvin Probe Force Microscopy (KPFM) must operate in single-pass KPFM mode; it must enable simultaneous operation with sSNOM imaging; - Tip Force Microscopy (TFM) must enable simultaneous measurement of local stiffness, adhesion, viscosity, and energy dissipation; - Piezo Force Microscopy (PFM) must enable measurement of local electromechanical material properties through the piezoelectric response; - conductive AFM must enable simultaneous measurement of topography and electric current between the tip and the sample in contact mode and must enable correlative measurements with s-SNOM; a dedicated sample stage for electromagnetic shielding of the sample must be provided. |
| 5.14 SUPPORTING EQUIPMENT | 1 | Probes for near-field and AFM-IR measurements The bidder shall offer probes compatible with the offered system for nano-IR microscopy, near-field mapping, and spectroscopy. The offered solution must meet at least the following requirements: - it must include at least 380 probes for near-field measurements; - it must include at least 380 probes for AFM-IR measurements; - the probes must be compatible with the offered system and with all associated measurement modes; - the probes must be suitable for routine mapping and spectroscopy. |
| 5.15 SUPPORTING EQUIPMENT | 1 | Licensed measurement software (1 licence) Licensed data processing software for at least 10 users. |
| 5.16 SUPPORTING EQUIPMENT | 1 | Additional equipment: FTIR system for automated microscopic imaging and mapping and point measurements The bidder shall offer an FTIR system for automated microscopic imaging and mapping and point measurements. The system must be a compact, standalone FTIR microscope with an integrated spectrometer and motorized functional units. It must enable automated measurements in ATR and reflection mode, as well as in transmission mode without additional user intervention or with an appropriate transmission unit included. All optical and measurement assemblies must be software-controlled. The system must enable chemical imaging using an IR imaging detector, as well as point measurements and mapping using a single-element detector. It must include at least: - an IR imaging detector with at least 1024 pixels; - an additional liquid-nitrogen-cooled MCT detector for point analysis and mapping; - beam splitter: ZnSe; - diode laser with at least a 10-year factory warranty; - the system must include an interferometer with a design enabling permanent alignment in order to ensure insensitivity to mirror tilt, mechanical vibrations, and thermal effects; - mechanical and automatic alignment of the interferometer shall not be acceptable; - 10-year factory warranty for the interferometer; - 5-year factory warranty for the source; - software-controlled switching between detectors; - all microscope components, including the source, laser, detectors, electronics, and interferometer, as well as the humidity level inside the instrument, must be continuously monitored; the user must be automatically informed via the software of any detected deviation; - the option of a second detector position or an equivalent solution. The spectral and measurement performance of the system must achieve at least: - spectral range for IR imaging of at least 5000 to 800 cm⁻¹, or wider; - spectral range of the MCT detector of at least 6000 to 700 cm⁻¹, or wider; - spectral resolution of 2 cm⁻¹ or better; - acquisition speed with the IR imaging detector of at least 500 spectra/s at a spectral resolution of 4 cm⁻¹ or better, at a lateral resolution of 5 µm; - acquisition speed with the MCT detector of at least 5 spectra/s at a spectral resolution of 4 cm⁻¹ or better; - wavenumber accuracy for the MCT detector of 0.05 cm⁻¹ or better; - wavenumber accuracy of the IR imaging detector for each pixel of 0.1 cm⁻¹ or better. The system must include an optical microscopy assembly with at least the following characteristics: - objective for reflection and ATR measurements with a numerical aperture of at least 0.6 or higher; - integrated motorized ATR crystal with integrated pressure control for automated ATR measurements; - switching between ATR measurements and visible observation of the sample must be software-controlled and automatic; - working distance of at least 30 mm or equivalent usability; - digital camera with at least 5 MP for sample inspection and selection of the measurement location; - visible illumination for sample inspection; - automatic switching between IR and visible mode; - automatic adjustment of the numerical aperture or an equivalent optical solution; - autofocus. The sample holder and stage movements must enable automated work with different types of samples. It must include at least: - motorized XY stage with a travel range of at least 75 mm × 50 mm; - minimum step size of at most 0.1 µm or less; - movement repeatability of 1 µm or better; - motorized Z travel with a range of at least 40 mm; - the option of control via software and via a dedicated controller or joystick, or an equivalent solution. The system must include at least the following additional equipment for extended usability: - transmission channel with motorized condenser and illumination; - motorized illuminated transparent knife-edge aperture with adjustment down to at least 5 µm × 5 µm or less; - sample holder for particle analysis on filters in transmission, suitable for at least three filters of 25 mm diameter; - sample holder for ATR microscopy of samples; - universal sample holder for microanalysis of films and fibers; - additional clamping elements for analysis of cross-sections and multilayer structures at a 90° angle. Software must be included in the offer and must enable system control, acquisition, processing, evaluation, visualization, and reporting. It must include at least: - software package for FTIR measurements and processing of spectral data; - software module for video inspection and for guiding the user through the measurement procedure; - software module for 3D data processing, evaluation, and visualization; - software module for control of the motorized stage and mapping; - the possibility of atmospheric compensation; - functions for normalization, baseline correction, derivation, spectral subtraction, averaging, and calculations; - spectral library search; - comparison of spectra and identification of materials; - quantitative evaluation; - automatic report generation; - instrument status monitoring; - support for OQ/PQ tests or equivalent performance verification. At least one software module for advanced identification of the chemical composition of samples based on reference spectral databases must be included, as well as at least one software module for automatic recognition, measurement, and statistical processing of particles, fibers, or other regions of interest in the microscopic image, or an equivalent software solution. Spectral libraries suitable for material identification must be included, at least: - a comprehensive ATR-FTIR library with a total of at least 26,000 spectra or more; - an additional specialized library for polymers, plastics, and additives with at least 1,000 spectra, or an equivalent specialized database. Additional software licences for at least 10 additional users in an academic environment, or equivalent, must be included. The offer must include a dedicated processing computer system for data acquisition and processing, at least with: * - Windows 11 Pro 64-bit operating system or an equivalent newer version; - at least an i7-class processor or better; - at least 128 GB RAM; - SSD with a capacity of at least 2 TB; - additional data drive with a capacity of at least 8 TB; - dedicated graphics card; - at least one widescreen display with a diagonal of at least 34"; - all necessary interfaces for operation of the system. *Special requirements for the dedicated processing computer and display under item 5.16 in accordance with the Regulation on Green Public Procurement are set out at the end of this document. |
| 5.17 SUPPORTING EQUIPMENT | 1 | Dedicated processing computer with the following specifications: * - MS Windows 11 Professional 64-bit operating system; - desktop/workstation computer with a processor with at least 16 physical cores and a maximum boost frequency of at least 5.7 GHz or better; - 32 GB RAM or more, 2 TB SSD or more; - 2 × 34" (or larger than 34") LED displays with Full HD resolution or higher than Full HD. * Special requirements for the dedicated processing computer and displays in accordance with the Regulation on Green Public Procurement are set out at the end of this document. |
| 5.18 | 1 | Basic user training Introduction to the equipment and basic training, to be provided by the bidder’s engineer in Slovenian or English, for up to four users, for a duration of at least 5 working days, 8 hours per day (including handover of all necessary data and passwords, as well as the travel and accommodation costs of the bidder’s engineer; the training shall be carried out at the location of FKKT UM designated by the contracting authority). This item shall be considered completed after delivery, installation, and commissioning of the equipment. |
| 5.19 | 1 | Advanced user training package The advanced training package shall comprise two separate advanced training sessions on site after acceptance (the training shall be carried out at the location of FKKT UM designated by the contracting authority) in Slovenian or English and shall include: - 2 advanced training sessions, each with a duration of 5 working days (8 hours per working day), no later than 31 days after completion of the basic training referred to in item 5.18. The advanced training must be provided for up to four persons and must include the travel, subsistence, and accommodation costs of the bidder’s personnel delivering the training. |
| 5.20 | 1 | Warranty (minimum 24 months of full coverage) The bidder must ensure at least 24 months of coverage from successful installation and successful completion of the acceptance test on site. - The coverage must include replacement of defective parts and all associated labour costs. - The coverage must include travel expenses and accommodation costs of the bidder’s personnel for all required on-site interventions. - The coverage must include all spare parts required to maintain the instrument in optimal condition and with high availability, at no additional cost during the coverage period. - The service package must include all consumable parts that need to be replaced according to service life/condition, together with on-site replacement performed by the bidder’s qualified personnel. - It must include at least one preventive visit per year, with a minimum on-site duration corresponding to two working days (8 hours per working day), including all travel and accommodation costs. - It must include an unlimited number of corrective service visits as needed, including urgent interventions, with labour, travel, accommodation, and required parts included. - It must include calibrations, testing, and performance verification required to maintain the instrument within specifications. - It must include unlimited remote diagnostics and support (e-mail and secure remote access/remote desktop). - The package must include free updates of the installed control and analysis software. - The package must include advanced user training in accordance with the offered service level. - The bidder must provide on-site service within one week and respond to e-mail notifications within 48 hours. |
| 5.21 | 1 | Remote access to the instrument must be ensured for the purposes of calibration and instrument control, at least for the duration of the warranty period (24 months). The instrument design must enable remote operation and fault diagnostics. |
| 5.22 | 1 | Delivery and installation on the client’s premises; 3rd floor at FKKT UM. Conditions for successful completion of the item and signature of the acceptance report: - delivery, installation, and commissioning of the equipment at the contracting authority’s site, performed by the bidder; - familiarization with the equipment and basic training referred to in item 5.18; - handover of the operating instructions and other technical documentation in digital form and in English; - the performance specifications (items 5.1 to 5.17) must be demonstrably fulfilled on site (FKKT UM) before signature of the acceptance documents. |

SPECIAL REQUIREMENTS FOR ITEMS 5.16 and 5.17

The equipment under items 5.16 and 5.17, Dedicated Processing Computer, must fully comply with the technical specifications, ensuring the fulfillment of the objectives in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 121/21, 132/23, and 43/25), namely:

DEDICATED PROCESSING COMPUTER (ITEMS 5.16 and 5.17)

Power Consumption MANDATORY: The personal computer must belong to the highest energy efficiency class available on the market. Method of proof: The bidder must attach the following to the tender: - a certificate showing that the product has a Type I environmental label or a label from another certification scheme (e.g. Energy Star V9.0, EPEAT, TCO) that meets the specified requirements. Appropriate supporting documents may also include alternative test results obtained by testing bodies accredited in accordance with ISO 17025, in compliance with IEC 62623.

ELECTRONIC DISPLAYS (ITEMS 5.16 and 5.17)

Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market, an energy class of at least D,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL).

The operation is co-financed by the European Union from the European Regional Development Fund under the European Cohesion Policy Programme 2021–2027, Policy Objective 1: A more competitive and smarter Europe by promoting innovative and smart economic transformation and regional ICT connectivity, under Priority 11: Development of Strategic Technologies for Europe – STEP, and under Specific Objective RSO1.6: Support for investments contributing to the objectives of the Strategic Technologies for Europe Platform (STEP).