Project name: GIGA Nuclear Magnetic Resonance – GIGA NMR

APPENDIX 2

DECLARATION OF COMPLIANCE WITH THE CONTRACTING AUTHORITY'S MINIMUM TECHNICAL REQUIREMENTS
AND SPECIFICATIONS OF THE OFFERED EQUIPMENT

| Project name: | GIGA Nuclear Magnetic Resonance - GIGA NMR |
| --- | --- |
| Public procurement No.: | 302/7-O-CATS/26 |
| Subject of public procurement: | Purchase and delivery of the Equipment for the Center for Solid-State Analysis |
| Lot: | Lot 2: X-ray photoelectron spectroscopy equipment for operation at near-ambient pressure (NAP-XPS) |
| Tenderer’s (bidder’s) name: |  |
| Tenderer’s (bidder’s) address |  |

| Name and type of equipment offered: | Name: | Type: |
| --- | --- | --- |
| Is the provider the manufacturer of the offered equipment? | YES ☐ | NO ☐ |
| If the answer is NO, state the manufacturer(s) of the offered equipment (name and address): |  |  |

INSTRUCTIONS The tenderer (bidder) shall complete and sign this appendix. In the “OFFERED” column, they shall indicate compliance with the requirement and, where necessary, specify the model, configuration, or proposed technical solution. In the “PROOF” column, they shall state the title of the supporting document, technical data sheet, brochure, catalogue, certificate, or provide a web link to the relevant technical specification (except where this is not provided for a specific item or where a declaration is required). The form has been prepared in accordance with the technical specifications for Lot 2 and also includes additional requirements for item 2.14 on green public procurement. The bidder shall load this Appendix 2 in PDF format and in Word format (if the form contains active hyperlinks to the required supporting documents), including attached evidence of compliance with the technical requirements, in the e-JN information system, in the section "Other attachments".

We hereby declare that the equipment offered for Lot 2 complies with the contracting authority's minimum technical requirements for the public procurement "Purchase of the Equipment for the Center for Solid-State Analysis" according to the specification set out below.

| Item No. | No. of pieces | REQUIRED | OFFERED | PROOF |
| --- | --- | --- | --- | --- |
| 2.1 | 1 | - Fully automated NAP-XPS base system capable of operation from UHV into the near‑ambient pressure regime, with demonstrated operation at pressures up to at least 50 mbar. - The instrument shall be a fully enclosed (housed), standalone laboratory NAP-XPS system suitable for multi-user operation. - The system shall require only standard laboratory utilities: electrical power, venting gas, compressed air (if required by the offered design), and Ethernet network connection for computer/control. - Sample introduction time (loading-to-measurement) of less than 2 minutes for near-ambient pressure measurement modes, without requiring pump-down of the analysis environment to UHV prior to initiating NAP measurements. - Capability to measure liquids, organic and porous samples, and to provide intrinsic charge compensation for semiconducting and insulating materials via background‑gas photoionization. - Electron spectrometer requirements: hemispherical electron analyser; one‑dimensional delay‑line detector with at least 400 channels (or more); differentially pumped lens system; automated in‑lens valve. Lens acceptance / collection efficiency: In the near-ambient pressure measurement configuration, the lens/analyser front-end shall provide an electron collection acceptance angle of ≥ ±20° (total ≥ 40°) or an equivalent collection solid angle, with performance demonstrated during acceptance testing. - The offered system shall achieve the following minimum performance on a polycrystalline, non-sputtered silver (Ag) reference sample, using monochromated Al Kα excitation at 40 W and 15 kV with an X-ray spot size of 300 μm (nominal), evaluated on the Ag 3d5/2 peak: a) UHV conditions: - At an energy resolution of FWHM = 0.85 eV ± 0.05 eV, the measured Ag 3d5/2 peak count rate shall be ≥ 45 kcps. - At an energy resolution of FWHM = 0.60 eV ± 0.05 eV, the measured Ag 3d5/2 peak count rate shall be ≥ 20 kcps. b) Near-ambient pressure conditions: - At 10 mbar N2, and at an energy resolution of FWHM = 1.0 eV ± 0.1 eV, the measured Ag 3d5/2 peak count rate shall be ≥ 6 kcps. c) Energy resolution (monochromated Al Kα): - The system shall be capable of achieving ≤ 0.60 eV FWHM on Ag 3d5/2 under UHV conditions. - X‑ray source requirements: monochromated Al Kα source; selectable spot size 200 µm to 1 mm in diameter, optimized to the analysis area. - Charge neutralization: the system shall provide intrinsic charge compensation via photoionization of the background gas, effective for insulating and semiconducting samples under near-ambient pressures (including ≥ 1 mbar). - Pumping system: pressure range from 1·10⁻7 mbar to >40 mbar; chemically resistant pumping architecture; turbomolecular pumps with oil‑free backing pumps. - In-situ cleaning: Include a downstream RF plasma cleaning capability (or equivalent) to support chamber/environment cleaning without compromising analytical components. - The offered NAP-XPS system shall enable analysis of vacuum-sensitive samples under controlled gas atmospheres and humid/wet environments at near-ambient pressures without requiring prior pump-down of the analysis environment to UHV as a precondition for starting the NAP measurement mode. For such measurements, the system design (automation, vacuum/gas handling, safety interlocks, and software workflows) shall support direct initiation of near-ambient pressure measurements after sample loading, consistent with safe operation of the offered instrument. - Gas dosing system: fully automated and software‑controlled with at least two gas dosers at the analysis position and with at least three mass‑flow controllers. - Imaging/navigation: at least three digital microscopes for sample navigation and documentation. - Automation and software: fully automated sample handling, vacuum Management and gas dosing; multiple user‑level interfaces; automated Reporting, chemical quantification, and assistance tools. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.2 | 1 | The system must include a quadrupole mass spectrometer with: - Residual gas analyser for UHV with mass range 1–200 atomic mass unit. - Detector system comprising a Faraday cup and a secondary electron multiplier. - Includes power supply and control/data acquisition software. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.3 | 1 | The bidder shall supply a room-temperature sample environment module suitable for near-ambient pressure XPS measurements and fully compatible with the offered NAP-XPS system. The module must meet the following minimum requirements: - Operating purpose: Enable experiments at room temperature under controlled atmospheres, including operation in the near-ambient pressure regime supported by the main instrument. - Positioning and alignment: Include a motorized 3-axis sample stage (X/Y/Z). - Sample capacity: Accommodate samples up to 120 mm diameter or larger and up to 40 mm height or higher without modification of the sample environment. - Addressable area: Provide an addressable sample area with a minimum diameter of 50 mm. - Electrical feedthroughs: Provide electrical contacts (or equivalent capability) for in-situ biasing, heating control interfaces, sensors, or an equivalent solution. - Observation and documentation: Include a digital microscope (or equivalent imaging system) enabling sample observation, navigation, and documentation during setup and/or measurement. - Integration: The module must be fully integrated with the system’s vacuum control, safety interlocks, and control software, enabling automated operation and reproducible workflows. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.4 | 1 | High-temperature near-ambient pressure sample environment with IR laser heating (HT-NAP sample environment) The bidder shall supply a high-temperature NAP sample environment, fully compatible with the offered NAP-XPS system, meeting the following: - Heating: IR laser heating (or equivalent) with demonstrated capability of ≥ 800 °C at 10 mbar N2. - Stage: motorized 3-axis sample stage (X/Y/Z). - Sample: Supports samples up to 120 mm in diameter or larger and 40 mm in height or higher. - Addressable area: ≥ 50 mm in diameter for the experiment without heating and ≥ 10 mm in diameter with holders for heating. - Interfaces: at least 5 electrical contacts (or equivalent) are needed. Sample holder: at least one sample holder for high-temperature tests for measurements at ≥ 800 °C must be included. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.5 | 1 | Thermoelectrically cooled sample holder (Peltier-type) for NAP operation (cooled sample holder) - The bidder shall supply a thermoelectrically cooled (Peltier-type) sample holder fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system. - The sample holder shall achieve sample temperatures of below −5 °C under an atmosphere of 10 mbar nitrogen. - The sample holder shall accommodate samples at least 10 mm in diameter or higher. - The sample holder shall include an integrated thermocouple (or equivalent temperature sensor) enabling closed-loop temperature measurement and control. The sample holder shall be compatible with a high-temperature sample environment supplied for the offered system. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.6 | 1 | Laser heated sample holder for NAP operation (high-temperature sample holder) - The bidder shall supply a laser heated sample holder fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system and the supplied high-temperature sample environment. - The sample holder shall enable sample temperatures above 800 °C under an atmosphere of 10 mbar nitrogen. - The sample holder shall include a sacrificial window (or equivalent protective element) positioned beneath the sample to protect the underlying components during high-temperature operation. - The sample holder shall accommodate samples at least 10 mm in diameter or higher. - The sample holder shall include an integrated thermocouple (or equivalent temperature sensor) enabling closed-loop temperature measurement and control. - • The sample holder shall be designed for operation in, and be fully compatible with, the high-temperature near-ambient pressure sample environment supplied with the system. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.7 | 1 | Inlet system for liquid vapours (liquid vapour dosing system) - The bidder shall supply an inlet system for dosing liquid vapours, fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system and its gas handling infrastructure. - The system shall include a mass-flow controller suitable for vapours with a flow range of 0.5–50 standard cubic centimetres per minute, or wider, but it needs to include 0.5–50 standard cubic centimetres per minute. - The system shall enable dosing of water vapour up to 20 mbar at the analysis position; capability to dose other liquid vapours shall be available upon request or by interchangeable liquid supply. - The system shall include a temperature-controlled liquid container/reservoir suitable for stable vapour generation and reproducible dosing. - The system shall provide a direct connection to the main gas inlet system of the instrument, with software-controlled operation where applicable. - The system shall include a bypass line to the fore-vacuum system enabling degassing of the liquid prior to dosing. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.8 | 1 | Electrochemical cell module for NAP operation (electrochemical interface module) - The bidder shall supply an in-situ refillable electrochemical cell module fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system for measurements at solid–liquid interfaces. - The module shall include an electrochemistry cell configured for three-electrode operation and shall provide liquid line connections (inlet/outlet) suitable for controlled electrolyte handling. - The supply shall include at least one peristaltic pump suitable for semi-static experiments (basic circulation or controlled replenishment). - The module shall include modular connection interfaces inside and outside the spectroscopy chamber, enabling (i) application-specific electrochemical cell variants and (ii) integration with an external/mobile wet-chemistry unit such as a reaction vessel with inert-gas purging and automated electrolyte dosing. - The module shall provide pressure control in the range of 2–1000 mbar or wider. The module shall be designed to allow connection of an external potentiostat, and the electrode/electrical interface shall be defined during implementation to ensure full operational compatibility. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.9 | 1 | Cluster ion source for sputter depth profiling under ultra-high vacuum (cluster ion sputtering module) - The bidder shall supply a gas cluster ion source fully compatible with the offered X-ray photoelectron spectroscopy system for sputter depth profiling under ultra-high vacuum conditions. - The system shall be an argon cluster ion beam source with an acceleration energy of 10 keV (or higher). - The system shall provide selectable cluster sizes. - The supply shall include an ion source comprising a cluster generation unit and an ionisation chamber. - The system shall include an ion-optical column incorporating a Wien mass filter (or equivalent) for selection/control of the cluster ion species. - The system shall include a differential pumping system suitable for maintaining ultra-high vacuum compatibility during operation. - The cluster ion source shall be fully integrated into the instrument control software, enabling automated depth profiling workflows, method setup, and data acquisition. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.10 | 1 | Vacuum transport suitcase for sample transfer - The bidder shall supply a vacuum transfer vessel fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system for transferring samples between the instrument and external vacuum systems and a glovebox. - The vacuum transport suitcase shall be capable of docking directly to the supplied sample environment of the system through a suitable vacuum interface, enabling controlled transfer without breaking vacuum at the docking point. - The vacuum transport suitcase shall be mechanically and dimensionally suitable for use inside a glovebox in Item 2.12, including the ability to be introduced into the glovebox and operated for sample loading/unloading under a protective atmosphere. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.11 | 1 | Sample vessel for protected-atmosphere transport (sample transport vessel) - The bidder shall supply a compact sample vessel suitable for transporting samples under a protective atmosphere and fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system. - The sample vessel shall provide a connection enabling docking to a loadlock and/or to the supplied vacuum transport suitcase for controlled transfer. The sample vessel shall support sample loading and sealing under a protective atmosphere. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.12 | 1 | Integrated glovebox module for NAP-XPS sample handling (glovebox module) - The bidder shall supply an integrated glovebox module fully compatible with the offered near-ambient pressure X-ray photoelectron spectroscopy system for inert-atmosphere sample handling and transfer. - The glovebox module shall be manufactured from stainless steel or an equivalent corrosion-resistant material suitable for long-term inert-atmosphere operation. - The glovebox module shall include an integrated gas purification/cleaning system suitable for maintaining low oxygen and moisture levels during operation. - The glovebox module shall maintain controlled residual oxygen and water concentrations below 1 part per million during normal operation (target specification). - The glovebox module shall provide pressure regulation enabling operation at ±15 mbar overpressure or underpressure relative to ambient pressure (or wider). The supply shall include all required mechanical and vacuum interface hardware/adapters enabling direct connection between the glovebox module and the NAP-XPS system for protected sample transfer. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.13 | 1 | Licensed software for measurement and data processing (at least one licence). Licensed data-processing software for at least 10 users. The control and analysis software shall provide full instrument control and data processing, including background subtraction, peak fitting, quantification, and automated reporting, for operation from UHV to at least 50 mbar. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.14 | 1 | Dedicated process computer with the following specifications: * - MS Windows 11 Professional 64-bit operating system, - desktop/workstation computer with a processor with at least 16 physical cores and a maximum boost frequency of at least 5.7 GHz, or better, - 32 GB RAM or more, 2 TB SSD or more, - 2 × 27" LED displays (27" or larger) with Full HD resolution or higher than Full HD resolution. * Note: Special requirements for item 2.14 in accordance with the Decree on Green Public Procurement are set out at the end of this document. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Document / URL: ________________ Reference / page: ________________ |
| 2.15 | 1 | Warranty (minimum 60 months total coverage). The bidder shall provide a minimum total coverage period of 60 months from successful installation and site acceptance test. - It shall cover replacement of defective parts and all corresponding labour costs. - It shall cover travel and living expenses of the bidder’s personnel for any required on-site intervention. - It shall cover all spare parts required to maintain the instrument in optimum condition and high availability, at no additional cost during the covered period. - The service package shall cover all consumables required according to lifetime/condition, including on-site replacement/exchange performed by the bidder’s expert personnel. - It shall include at least one preventive maintenance visit per year with a minimum on-site duration equivalent to two working days (8 hours per working day), including all travel and accommodation costs. - It shall include unlimited corrective service visits as required, including emergency visits, with labour, travel, accommodation, and necessary parts included. - It shall include calibration, testing, and performance verification activities required to keep the instrument within specification. - It shall include unlimited remote diagnostics and support (e-mail and secure remote access/remote desktop). - The service package shall include free software updates for the installed control and data analysis software. - The service package shall include extended customer training consistent with the offered service level. - The bidder must provide on-site service within one week unless otherwise agreed, and must respond to email notifications within 48 hours. | We declare that we offer the entire item No. 2.15 YES ☐ NO ☐ (Contracting Authority’s note: for the item in question, the above statement shall be deemed as evidence of compliance with the Contracting Authority’s requirement). |  |
| 2.16 | 1 | Basic operation training Acquaintance with the equipment and basic training, provided by the contractor’s engineer in Slovene or English, for up to four users, lasting at least 5 working days, 8 hours per day (including the handover of all necessary data and passwords, as well as the travel and accommodation costs of the contractor’s engineer; the training shall be conducted at the FKKT UM location designated by the contracting authority). This item shall be deemed completed upon delivery, installation, and commissioning of the equipment. | We declare that we offer the entire item No. 2.16 YES ☐ NO ☐ (Contracting Authority’s note: for the item in question, the above statement shall be deemed as evidence of compliance with the Contracting Authority’s requirement). |  |
| 2.17 | 1 | An advanced user training package The advanced training package comprises one separate on-site advanced training session after acceptance (the training shall be conducted at the FKKT UM location designated by the contracting authority) in Slovene or English, and includes: Advanced training lasting 5 working days (8 hours per working day), to be conducted no later than 31 days after completion of the basic training under item 2.16. The advanced training must be delivered for up to four persons and must include the travel, living, and accommodation costs of the contractor’s training personnel.. | We declare that we offer the entire item No. 2.17 YES ☐ NO ☐ (Contracting Authority’s note: for the item in question, the above statement shall be deemed as evidence of compliance with the Contracting Authority’s requirement). |  |
| 2.18 | 1 | Provided remote access to the device for the needs of calibration and control of the device, at least for the duration of the warranty period (60 months). The instrument design must facilitate remote operation and fault diagnostics. | We declare that we offer the entire item No. 2.18 YES ☐ NO ☐ (Contracting Authority’s note: for the item in question, the above statement shall be deemed as evidence of compliance with the Contracting Authority’s requirement). |  |
| 2.19 | 1 | Delivery and installation on the client’s premises; 3rd floor at FKKT UM. Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation, and commissioning of the equipment on the client’s premises by the bidder, - familiarization with the equipment and basic training under item 2.16, - handover of the operating instructions and other technical documentation in digital form and in English, - performance specifications (items from 2.1 to 2.14) must be demonstrated on-site (FKKT UM) before signing acceptance documents. | We declare that we offer the entire item No. 2.19 YES ☐ NO ☐ (Contracting Authority’s note: for the item in question, the above statement shall be deemed as evidence of compliance with the Contracting Authority’s requirement). |  |

ADDITIONAL REQUIREMENTS FOR ITEM 2.14 - GREEN PUBLIC PROCUREMENT

| AREA | REQUIRED | OFFERED | PROOF |
| --- | --- | --- | --- |
| Dedicated processing computer – item 2.14 | Power Consumption MANDATORY: The personal computer must belong to the highest energy efficiency class available on the market. Method of proof: The bidder must attach the following to the tender: - a certificate showing that the product has a Type I environmental label or a label from another certification scheme (e.g. Energy Star V9.0, EPEAT, TCO) that meets the specified requirements. Appropriate supporting documents may also include alternative test results obtained by testing bodies accredited in accordance with ISO 17025, in compliance with IEC 62623. | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Indicate the evidence of compliance with the energy efficiency requirement! Document / URL: ________________ Reference / page: ________________ |
| Electronic displays – item 2.14 | Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market, an energy class of at least B,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL). | YES ☐ NO ☐ Model / configuration: ________________ Offered specification: ________________ | Indicate the evidence of compliance with the energy efficiency requirement! Document / URL: ________________ Reference / page: ________________ |

| Place and date: | Seal: if the bidder operates using an official company seal | Signature of authorized person: |
| --- | --- | --- |

The operation is co-financed by the European Union from the European Regional Development Fund under the European Cohesion Policy Programme 2021–2027, Policy Objective 1: A more competitive and smarter Europe by promoting innovative and smart economic transformation and regional ICT connectivity, under Priority 11: Development of Strategic Technologies for Europe – STEP, and under Specific Objective RSO1.6: Support for investments contributing to the objectives of the Strategic Technologies for Europe Platform (STEP).