## Summary
Project |  | Veterinary Equipment for the New Veterinary Faculty Building in Ljubljana |  | Contracting Authority | University of Ljubljana
Type of Works |  | EQUIPMENT |  | End User | New Veterinary Faculty
Content |  | EQUIPMENT SCHEDULE |  | Date | June 2026
| 1.4.3. | EQUIPMENT SPECIFICATIONS
SUMMARY OF EQUIPMENT COSTS
Item No. | Code | Description | Total Price (excluding VAT) in EUR
| 1 | Equine Recovery Box Wall and Floor Covering System | 0
|  | Contingency Works (5%) | 0
|  | SUBTOTAL | 0
|  | VAT 22% | 0
|  | TOTAL INCLUDING VAT | 0
| NT.H.1 | kombinirani, cca +2°C do +15°C, -10°C do -24°C , cca 250l+250l
| NT.H.2 | kombinirani, cca +2°C do +15°C, -10°C do -24°C , cca 170l+170l
| NT.H.3 | cca -10°C do -24°C , cca 600l
| NT.H.4 | cca -10°C do -24°C , cca 350l
| NT.H.5 | podpultni, cca +2°C do +15°C, cca 100l
| NT.H.5.1 | podpultni, cca +2°C do +15°C, cca 100l, z vtičnico znotraj
| NT.H.6 | cca -80°C , cca 500l
| NT.H.7 | cca +2°C do +15°C, cca 1500l
| NT.H.8 | podpultni, cca -16°C do -28°C, cca 70l
| NT.H.9 | cca +2°C do +15°C, cca 350l
| NT.H.10 | cca -80°C , cca 1000l
| NT.H.11 | cca -80°C , cca 800l
| NT.H.12 | cca +2°C do +15°C, cca 350l, steklena vrata
[10 empty] | Izvlečna prečka za obešanje; | 01.22.72 | MARA STOJEČA, GARDEROBNA, S POLNIMI VRATI, LESENE POLICE - 110/40/215 cm
[6 empty] | Na tleh; | Fiksirana na steno (1.95 m nad tlemi); | 1 x dvokrilna, polna; | 5 x polica lesena;

## Splošno
General requirements for the supply and installation of equine recovery box coverings:
The selected Contractor shall, upon request of the Contracting Authority following contract award, submit the following certificates and supporting documentation demonstrating compliance of the offered system:
1. | General compliance with quality management and organization requirements for manufacturing, supply and installation | ISO 9001
ISO 14001
2. | For all offered system components, compliance with applicable EU legislation (issued by an independent accredited institution authorized to issue such certificates within the EU) | Declaration of Conformity (Certificate of Conformity), CE marking
3. | Resistance of the offered materials to H₂O₂ disinfection | Certificate confirming material resistance to H₂O₂ or test report
4. | Compliance with REACH Regulation (EC) No. 1907/2006 | Declaration
5. | Manufacturer’s confirmation of representation/service support for the offered system in the Republic of Slovenia (for foreign suppliers) | Declaration
The following conditions shall be included in the offered price:
1. The offered system shall have an expected service life of at least 10 years, including the availability of spare parts and replacement components.
2. The warranty period shall be 2 years from the date of final acceptance by the Contracting Authority.
3. The Contractor shall provide Operating and Maintenance Manuals.
4. Upon completion of installation, the Contractor shall provide user training for the safe operation and maintenance of the system. The training shall be documented by a signed training record.
5. During the warranty period, the maximum response time for corrective intervention shall be 2 working days from notification of the defect or failure.
6. The maximum repair time for defects or failures shall be 14 calendar days from the date of notification.

## Equine Recovery Box Wall and Fl
Project |  | Veterinary Equipment for the New Veterinary Faculty Building in Ljubljana |  |  | Contracting Authority | University of Ljubljana
Type of Works |  | EQUIPMENT |  |  | End User | New Veterinary Faculty
Content |  | EQUIPMENT SCHEDULE |  |  | Date | June 2026
| 1.4.3. | EQUIPMENT SPECIFICATIONS
Equine Recovery Box Wall and Floor Covering System
Item No. | Code | Description | Area (m²) | Unit Price per m² (excluding VAT) in EUR | Total Price (EUR, excl. VAT)
GENERAL DESCRIPTION OF WORKS
Supply and installation of a special protective covering system for three (3) equine recovery boxes.
Quantities:
* wall and door cladding: approximately 130 m²;
* flooring system: approximately 50 m².
Final quantities shall be determined based on site measurements prior to execution of the works.
THE CONTRACT PRICE SHALL INCLUDE:
* supply of all materials;
* transportation;
* unloading;
* site measurements;
* preparation of shop drawings and installation drawings;
* substrate preparation;
* installation and commissioning;
* all fixing and fastening elements;
* all finishing and trim elements;
* sealing of all joints and interfaces;
* treatment of interfaces between floors, walls and doors;
* treatment of ventilation ducts, grilles and openings;
* reinforcements around ventilation ducts and openings;
* cleaning upon completion of the works;
* removal and disposal of packaging materials;
* handover of the completed system to the Contracting Authority.
The Tenderer shall submit technical data sheets, a Declaration of Conformity, evidence of compliance with REACH Regulation (EC) No. 1907/2006, and operation, cleaning and maintenance instructions for the proposed system.
In the area of ventilation ducts, ventilation openings, grilles, diffusers and other built-in elements, the Contractor shall provide adequate mechanical reinforcement of the wall cladding system.
The reinforcements shall be manufactured from stainless steel grade AISI 304 or an equivalent material and shall provide:
* protection of ventilation components against mechanical damage;
* protection of the wall cladding system against deformation and tearing;
* resistance to impacts and loads resulting from horse movement and falls during recovery;
* watertightness of all joints, interfaces and penetrations;
* suitability for cleaning and disinfection.
All finishes around ventilation openings shall be free from sharp edges, protruding parts or other elements that could cause injury to animals.
The Contract Price shall include:
* all required cut-outs;
* all steel reinforcements;
* all finishing profiles;
* all fixing and fastening elements;
* sealing of all joints and interfaces;
* all adaptations required to suit the actual site conditions.
| OB5 | WALL AND DOOR CLADDING | 130 |  | 0
Supply and installation of a multilayer protective wall and door cladding system for equine recovery boxes. The system shall be installed on reinforced concrete walls and door leaves of the recovery boxes and shall provide protection against injuries caused by impacts during recovery from general anaesthesia.
The system shall include, at a minimum:
* a water-resistant backing board with a minimum thickness of 12 mm;
* an elastic impact-absorbing cellular rubber layer with a minimum thickness of 30 mm;
* reinforcement in the plinth area to prevent separation between the flooring and wall systems;
* a wear-resistant polyurethane (PU) top layer with a minimum thickness of 4 mm;
* watertight sealing of all joints between individual elements;
* mechanical fixing to the reinforced concrete substrate.
The wall cladding system shall provide:
* resistance to horse impacts;
* abrasion resistance;
* resistance to urine and cleaning agents;
* watertightness;
* ease of cleaning and disinfection.
Cladding height: up to approximately 2.40 m above finished floor level (FFL).
The item shall include all necessary modifications, cut-outs, finishes and reinforcements around:
* ventilation ducts;
* ventilation grilles;
* ventilation openings;
* door frames;
* penetrations and other built-in elements.
Upon completion, no sharp edges, exposed fixing elements or unsealed joints shall remain visible.
Reference system: Paul Gröne AWB Walls Horse Protect VETMED 46 or equivalent.
| FT5.3 | FLOORING SYSTEM | 52 |  | 0
Supply and installation of a multilayer shock-absorbing flooring system for equine recovery boxes.
The system shall include, at a minimum:
* a cellular rubber layer with a minimum thickness of 50 mm;
* full-surface bonding without voids or air pockets;
* a perimeter sealing and stabilisation strip made of a liquid elastomeric system, with a minimum height of 50 mm and a minimum width of 250 mm;
* a homogeneous solid rubber floor covering with a minimum thickness of 15 mm;
* watertight sealing of all joints;
* full-surface contact bonding.
The total thickness of the flooring system shall be not less than 65 mm.
The flooring system shall be resistant to:
* hoof impacts;
* mechanical loads generated by horses during recovery;
* urine;
* cleaning agents;
* disinfectants;
* wet cleaning procedures.
Reference system: Paul Gröne AWB Flooring Horse Protect VETMED 65 or equivalent.
0