Project name: Package 23 “Small-angle X-ray scattering (SAXS) beamline at Elettra-Sincrotrone Trieste”

Declaration of Compliance with the Contracting Authority’s Minimum Technical Requirements and Specifications of the Offered Equipment

Tenderer’s name:

Tenderer’s address:

Subject-matter of the public procurement: “Hybrid photon detector for use on a small-angle X-ray scattering (SAXS) beamline”

Designation of the public procurement: 302/6 – HF-DET/26

| Is the tenderer the manufacturer of the offered equipment? | YES ☐ | NO ☐* |
| --- | --- | --- |
| *If your answer is NO, state the manufacturer/manufacturers of the offered equipment (name, address): |  |  |

INSTRUCTION:

- This Attachment No. 2 has to be completed and signed. The tenderer must mark whether the offered equipment is meeting the required technical specifications of the offered equipment according to all Contracting Authority’s requirements. In the column “MEANS OF PROOF” the tenderer must state the name of the attached means of proof. The tenderer may also provide a link to the website of the equipment manufacturer which includes an indication of the relevant technical specifications.

- The tenderer uploads this Attachment No. 2 in PDF format or also in Word format, if the Form includes active web links to the means of proof, to the e-JN information system, section “Other Attachments”.

We declare that the offered equipment fully complies with the minimum technical requirements of the Contracting Authority for the subject-matter of the public procurement with the title “Procurement of an in-vacuum undulator system for integration into a synchrotron beamline” and that we are prepared to perform the public contract according to the following specification:

| EQUIPMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIRED | OFFERED |  | MEANS OF PROOF |
| 1.1 | Scope of Supply |  |  |  |
|  | The tenderer shall supply one (1) complete, fully functional hybrid photon counting (HPC) X-ray area detector system for use in synchrotron beamlines, including at least: - detector head, - detector control unit/acquisition server, - all required power cables, fiber links, interface cables, and connectors, - control software and documented programmatic interface for external control and data acquisition, - mechanical mounting interface hardware required for safe installation, - all documentation necessary for operation, integration, safety, and maintenance, - installation, commissioning, functional verification, and user training at the Contracting Authority’s site. Detector Technology - The detector shall be a hybrid photon counting (single-photon counting) X-ray area detector suitable for continuous high-rate operation use with synchrotron light sources. - The detector shall be capable for future vacuum compatibility. - The detector shall allow double gating mode. Geometry and Imaging Area - Pixel size shall be 150 µm × 150 µm or smaller. - Total pixel matrix shall be at least 1550 × 1600 pixels. - Active imaging area shall be at least 233.0 mm × 244.0 mm or larger. - The detector may be modular; any module gaps shall be fully specified, and stable calibration/flat-field correction capability shall be provided. - The fraction of defective pixels shall be less than 0.1%. Sensor and Energy Performance - Sensor material shall be silicon. - Sensor thickness shall be 450 µm or equivalent ensuring comparable performance across the specified energy range. - Usable X-ray energy range shall be at least 6–40 keV. - Adjustable energy threshold range shall be at least 4–30 keV. - The detector shall provide four (4) independent energy thresholds. Speed, Dead Time, and Count Rate - The detector shall support continuous acquisition. - Readout dead time shall be 100 ns or less. - Maximum sustainable count rate shall be at least 1 × 10⁷ photons/s/pixel under manufacturer-specified conditions. - Continuous frame rate shall be at least 1900 Hz in 8-bit mode. - Continuous frame rate shall be at least 950 Hz in 16-bit mode. - A region-of-interest (ROI) acquisition mode shall be available with a frame rate of at least: - 4000 Hz in 8-bit mode and - 2000 Hz in 16-bit mode. - Any frame-rate dependence on the number of active thresholds shall be stated and shall not reduce performance below the minimum requirements stated above when operated in the corresponding acquisition configuration. Data Output and Bit Depth - The system shall support image output in 8-bit and 16-bit modes as a minimum. - The system shall provide metadata recording for acquisition parameters, including thresholds, timing, and operating settings. - Time-stamped acquisition shall be supported. Synchronization and Triggering - The system shall provide external trigger input and trigger output for synchronization with external instrumentation. - Trigger electrical specifications and absolute maximum ratings shall be provided by the tenderer and shall be compatible with typical laboratory control logic. Control Unit / Acquisition Server - A dedicated control unit shall be supplied as part of the system. - The control unit shall support continuous operation and provide stable detector control and data handling. - Optical connectivity between the detector head and the control unit shall be provided using single-mode fiber links with appropriate transceivers and connectors. - The control unit shall provide Ethernet connectivity for control and data access. - The system shall provide a documented network-accessible control interface suitable for integration into higher-level instrument control software. - The system shall provide the control through REST based API. Power Requirements - Detector head power input shall be compatible with 100–240 VAC, 50/60 Hz. Cooling and Thermal Stabilization - The detector system shall require and support an external thermal stabilization unit for operation. - Normal operating coolant setpoint shall be 23 °C. - Maximum permissible coolant pressure shall be 3 bar or less. - The system shall specify coolant composition requirements and corrosion-inhibitor compatibility constraints. - All necessary coolant connections for integration shall be provided. Dry Gas Purge / Humidity Protection - The detector shall support continuous purge with dry air or nitrogen. - Purge gas shall be oil-free. - Purge gas inlet pressure requirement shall be compatible with 1–2 bar. - Purge gas flow requirement shall be compatible with 5–10 L/h. - The system shall specify humidity limits and recommended operating humidity for performance. Environmental Operating Limits - Operation shall be supported within an ambient temperature range of 20–35 °C. - Operation shall be supported under non-condensing humidity conditions consistent with stable detector performance. - The system shall provide internal monitoring and protective behavior for out-of-range temperature and humidity conditions. Mechanical Integration - Detector dimensions and mass shall be provided by the tenderer. - The detector shall provide a defined mounting interface using threaded attachment points. - Maximum permissible screw-in depth, clearances, and ventilation clearance requirements shall be specified and shall enable safe installation. - The tenderer shall provide dimensional drawings sufficient for integration. | ☐ YES | ☐ NO |  |
| 1.2 | Warranty and Service |  |  |  |
|  | - Warranty for the complete delivered system shall be at least 12 months. - The tenderer shall provide service procedures, support contact channels, and spare parts availability information. | ☐ YES | ☐ NO |  |

| Place and date: | Stamp: | Signature of the responsible person: |
| --- | --- | --- |

INSTRUCTION:

- The tenderer fills in the required fields, provides the signature and stamp.

- The Tenderer uploads the Form “Attachment 2: Declaration of Compliance with the Contracting Authority’s Minimum Technical Requirements and Specifications of the Offered Equipment” at the time of submitting the tender to the e-JN information system, section “Other Attachments”.

The project No. P23-79 with the title “Small-angle X-ray scattering (SAXS) beamline at Elettra-Sincrotrone Trieste” obtained under the Public Call for (co-)Funding the Purchase of Research Equipment (Package 23, No. 6316-9/2024-4) is co-funded by the Slovenian Research and Innovation Agency.

The tenderer must mark whether the offered equipment is meeting the required technical specifications of the offered equipment according to all Contracting Authority’s requirements.

The tenderer must state the name of the attached means of proof. The tenderer may also provide a link to the website of the equipment manufacturer which includes an indication of the relevant technical specifications.